## Slide 1: The HYSPEC Polarized Beam Spectrometer
Barry Winn, Mark Hagen, Mark Lumsden, Melissa Graves-Brook David Anderson, Xin Tony Tong
B2.5, rm G, 3:15-3:30 PM

## Slide 2
HYSPEC History and People
Mark Barry Tony David
Hagen Winn Tong Anderson
Melissa
Graves-Brook
Steve Shapiro
Igor Zaliznyak
Bill Leonhardt
Initial Proposal:
Principal Investigators &
BNL Engineer:
IDT Executive also includes:
At ORNL:
Instrument construction complete
Commissioning with unpolarized neutrons mostly done Polarized neutron commissioning beginning
R. McQueeny J. Rhyne M. Kenzelmann

## Slide 3: Hybrid Spectrometer: a cross between
A direct geometry spectrometer that selects Ei and E resoultion, and measures variable Ef of scattered neutrons using ToF

## Slide 4: Hybrid Spectrometer: a cross between
…and a triple-axis spectrometer’s vertical focusing array and a variable direction final flight path
60o x 15o r = 4.5 m Ar flight path

## Slide 5: The focusing element focuses neutrons from a vertical trumpet guide system
15 cm high guide, PG focusing array focused at sample with ~2 cm high spot size
Gold foil measurement at sample position, 1.8 m between focus element and sample, Ei=15 meV, Fermi Chopper frequency 180 Hz : 4.2E5 c/s/MW/cm2
Vanadium incoherent isotropic scatter to detector array at 70o
Chart: Vanadium 6.4 mm dia rod
| 6.04 | 15.12 | 19.82 | 31.16 | 40.58 | 61.9 | 83.41 | 93.8
 | 271.5982 | 1068.1375 | 880.1078 | 387.9432 | 297.6363 | 193.5234 | 145.1575 | 121.2226
Pyrolitic
Graphite
3.8 meV < Ei < 95 meV

## Slide 6: And we have the option to polarize the incident beam with a Heusler crystal array
Energy(meV) | Flipping Ratio of Heusler and guide fields pre-sample
50 | 23.01
14.59 | 25.44
7.5 | 28.41
Heusler
(polarizing)

## Slide 7: Flux / Resolution tradeoff using Fermi Chopper Frequency
V rod 6.4 mm diameter
Chart: Elastic line width
| 30 | 60 | 90 | 120 | 150 | 180
Ei=3.8 meV | 0.12391 | 0.073831 | 0.052304 | 0.041634 | 0.034356 | 0.032012
Ei=7 meV | 0.19908 | 0.11716 | 0.045321 | 0.033907 | 0.023164
Ei=20 meV | 0.20841 | 0.1581 | 0.063481 | 0.041922 | 0.033537
Ei=84.8 meV | 0.16794 | 0.12706 | 0.10143 | 0.081335 | 0.069308
Chart: Flux tradeoff with Fermi Frequency
| 30 | 60 | 90 | 120 | 150 | 180
Ei=3.8 meV | 5.0143 | 2.8589 | 1.9434 | 1.4699 | 1.1868 | 1
Ei=7 meV | 5.1249 | 2.923 | 1 | 0.59974 | 0.43394
Ei=20 meV | 5.04 | 2.8336 | 1 | 0.63431 | 0.44401
Ei = 84.8 meV | 1 | 0.80856 | 0.69903 | 0.62709 | 0.57187
Straight-blade Fermi chopper,
30-420 Hz

## Slide 8: Useful to study phonon excitations on single crystals
FeSb2, Ei=51.58 meV, Fermi chopper 420 Hz, High Detector Vessel angle, Multiple sample q, 30 min/angle, Mantid to NxSPE, Dave Mslice: T=3, 100, 300 K
I. Zaliznyak (BNL), A. Savici, A. Christianson, B. Winn, M. Hagen, R. Hu and C. Petrovic (in preparation)

## Slide 9: Or magnetic excitations in single crystals or powders
Very dispersive magnetic excitations in the ground state of the quantum magnet La(2-x)Ba(x)CuO4 with x=0.025 can be seen emanating out of ~ 1/2 1/2 L. These particular data, taken with HYSPEC at T=4 K, integrate over much of the direction normal to L, that is normal to the two dimensional copper-oxide planes in this layered magnet. These spin excitations are gapless and are known to extend to very high energies (~ 200 meV). Acoustic phonons can also be seen emanating out of the nuclear zone centers near 11L.
K. Fritsch, B. Gaulin (McMaster U), in preparation

## Slide 10: Getting ready for your samples
Dedicated
Cryostat
Dedicated closed cycle refrigerator compatible with 3He polarization analysis
Spare CCR
Larger OVC’s and new heat shields for both CCR’s

## Slide 11: Minimizing background…
70% already removed by shielding BL 10 near monolith
One frame shown (1/60 sec), with HYSPEC shutters closed
Leading edge coincides with subsequent spallation events
Tail is epi-cadmium (>500 meV)
Ei not yet useful at 4 meV, 9 meV and 27-38 meV
STILL HUNTING!

## Slide 12: And preparing to commission for polarization analysis
Status: See T Tong’s talk, Wednesday at 11:00 AM, B4.2, and D. Brown’s poster Tuesday at BP2.8
Option 1: 3He polarization analyzer
and Filling Station
Option 2: Polarizing supermirror
array from PSI

## Slide 13: Summary
Unpolarized commissioning mostly complete
Polarized commissioning has begun
  Through spring 2013
Limited availability in user program in spring 2013
  Unpolarized neutrons only
Expanding Instrument Team at ORNL

## Slide 14: It takes a village
Engineering & Design
  J. Terrell, R. Huerto
Installation
  R. Connatser, S. Proffitt
  D. Engle, W. Dawkins, C. Fletcher, S. Vasques, J. Brackett
  HAZELWOOD
Electrical & Controls
  J. Moss, R. Saethre
  A. Groff, R. Morgan
  Research Mechanics
Interlocks & Safety
  P. Wright, J. Proulx
  G. Rowland, P. Abston
Vacuum
  J. Price, C. Stone, R. Morton
Choppers
  W. McHargue, J. Garrett
DAS
  A. Parizzi, T. Thompson, G. Greene, M. Ruiz-Rodrigues
  M. Yao, J. Kohl, M. Sundaram, P. Zivanovic
Data Reduction & Analysis
  S. Campbell, A. Savici