Options for wide angle TOF neutron polarization analysis with polarized 3He

L. Passell1, L.D. Cooley1, W.C. Chen3, T. R. Gentile 3,V. Ghosh1, M. Hagen 2, W.T. Lee2, W. Leonhardt1, S.M. Shapiro1, I. Zaliznyak1

1Department of Condensed Matter Physics and Material Science, Brookhaven National Laboratory, Upton, NY 11973-5000, USA

2Spallation Neutron Source, Oak Ridge National Laboratory, Oak Ridge, TN 37830, USA

3National Institute of Standards and Technology (NIST), Gaithersburg, Maryland 20899, USA

The large and strongly spin-dependent thermal neutron absorption cross section of  3He [σ(I-½)>>σ(I+½)] has the potential to make polarized 3He gas an efficient and versatile neutron polarizer and polarization analyzer.  And indeed, with the recent development of optical pumping techniques capable of producing 3He gas polarizations of 70-75 percent, 3He polarization analyzers are starting to appear in neutron spectroscopic applications.  Thus far, however, their use has been limited by a number of  incompletely resolved problems.   Of these, the most difficult by far is the extreme sensitivity of the gas polarization to stray magnetic fields.  To obtain a satisfactory 3He gas polarization lifetime (τ1/2 greater than, say, 100 hours) in the presence of a significant sample field, experience shows that the spatial uniformity (ΔH/H) of the holding magnetic field that defines the direction of nuclear polarization in the cell needs to be on the order of  4-6x10-4/cm!   There are also other impediments to the use of  3He. Among these are the sensitivity of the gas polarization lifetime to the choice of gas cell wall material, beam-related background generated by scattering in cell entrance and exit windows, limitations imposed on cell volumes and on gas pressures by neutron optics constraints on the shape of the cells and the handling problems which arise because the cells are optically pumped off-line and must be periodically interchanged as the gas polarization decreases to maintain a reasonable working efficiency.   Here we examine some of these problems in the context of using 3He polarization analyzers on neutron spectrometers with wide angular acceptance deteector arrays such as those typically employed in time-of-flight spectrometers.

In wide-angular-acceptance applications, uniform efficiency of polarization analysis requires a "wrap-around" cell geometry (such as that in shown schematically Fig. 1) so that all sample-to-detector flight paths pass through the same amount of polarized 3He gas.   Typically such cells will be 10-15 cm long, operate with gas pressures in the range of 1-2 atmospheres and have holding fields ranging from a few gauss to 100 gauss or more.  We consider below how such a cell might be incorporated into three representative wide-angular-acceptance, low temperature sample environments:  (i) a low field (<0.1T) environment employing either a closed-cycle refrigerator or an ILL top-loading cryostat, (ii) a medium field (1-3T) environment utilizing either a conventional electromagnet or a cryogen-free superconducting magnet with magnetic screening of the fringe fields done by conventional shielding materials and (iii) a high field (10-15T) environment in which the sample field is generated by a compensated superconducting magnet and with both magnetic screening and the cell holding field provided by a passive, persistent-mode superconducting shroud.   As will be evident, as the sample field increases, the cell holding field problem becomes progressively more difficult and complex to resolve.

(i) Small sample field (Hs<0.1T)

Figure 1

When the sample field is small it is feasible to position the 3He cell and sample inside a pair of Helmoltz coils as shown in Figure 1.  In this case both the cell and sample fields are produced by the Helmholtz coil pair, the sample field being simply a guide field to maintain the polarization of the incident and scattered neutron beams.  Computer simulations show that to obtain the necessary holding field uniformity the cell should be located at the center of symmetry of the Helmholtz pair and be of small volume, the latter condition also implies that a wide-angular-acceptance  cell needs to be as close to the outermost sample radiation shield as possible.  The simulations also show that air-cooled copper coils will produce fields of the order of 0.01T of satisfactory uniformity over a volume that can encompass a cell with an angular acceptance of about 60 degrees.  Although not shown in the figure, an arrangement of three orthogonal Helmholtz coil pairs (the so-called PASTIS coil geometry(1) ) is also a possibility.  This would add to experimental versatility by permitting the sample and cell field to be aligned in any direction with respect to the scattering plane.

(ii) Intermediate sample field (Hs ~1-3T)

Figure 2

Sample fields in the 1-3 T range can be supplied by either conventional electromagnets or by cryogen-free superconducting magnets.   But even with the most careful compensation such magnets are likely to produce enough stray field to significantly distort the 3He cell holding field which - as shown schematically in Figure 2 - will in this case have to be supplied by a separate solenoid, preferably one with aluminum windings thin enough to be reasonably transparent to neutrons.   To obtain a workable gas polarization lifetime it will therefore almost certainly be necessary to not only move the cell away from the sample thermal radiation shield but screen the cell solenoid with a magnetic shield with entrance and exit windows for the scattered neutrons.  The increased sample-to-cell distance and the need for magnetic shielding will have several non-trivial consequences.  One is that to maintain the angular acceptance of the system the lateral and vertical dimensions of the cell will need to be increased in proportion to the increased sample-to-cell distance.  The other is that conventional magnetic shielding is not very transparent to thermal neutrons and, as Table I shows, is only effective in fields less than ~ 0.1T.  Assuming the maximum fringe field condition can be satisfied at an acceptable distance from the sample magnet, it remains an open question whether a magnetic shield with both neutron entrance and exit windows will screen the stray field enough to make it possible to achieve the needed holding field uniformity.

Table I

Material                	Field that can be shielded

(Gauss)

Annealed μ metal	       	    ~10

Partly-annealed μ metal                   ~100

Fe0.97Si0.03			 ~1000

(ii) High sample field (Hs ~10-15T)

A typical compensated, 15T split-pair superconducting sample magnet will produce a stray field of roughly 0.5-1T at distances of  ½ meter (which is where a. 3He cell of reasonable size would have to be located).  Given that conventional magnetic shielding is unlikely to be effective in a field of this magnitude, the only realistic alternative is to think in terms of putting the cell inside a hollow-center, ILL CRYOPOL-type dewar(2) containing a passive, persistent-mode superconducting sleeve that would simultaneously supply a cell holding field of satisfactory uniformity and screen the stray field from the sample magnet.  In such an application it is imperative that the superconducting sleeve, dewar walls and internal radiation shields on either side of the sleeve be thin enough to transmit neutrons without excessive attenuation and that the sleeve have sufficient current carrying capacity to supply the necessary magnetic screening.  Table II lists the superconducting materials that could, at least in principal, beemployed.

Table II

Material                	Field that can be shielded

with 1 mm thick foil

(Tesla)

Nb	       	    	~0.2

Al clad Nb0.37Ti0.63                                      ~1.0

Al clad Nb3Sn		                  ?

There are established techniques for forming persistent-mode, superconducting Nb sleeves (such as that used in the CRYOPOL instrument) but the current carrying capacity of Nb is probably too low to screen the fringe field from a high-field sample magnet.   On paper, Al clad Nb0.37Ti0.63  would appear to be an attractive alternative because it has the necessary  current carrying capacity and good mechanical properties as well. But Al clad Nb0.37Ti0.63 consists of alternate layers of Nb0.37Ti0.63 wire mesh and thin Al foil and it has yet to be demonstrated that a  monolithic, persistent-mode sleeve could be fabricated from this material.  Al clad Nb3Sn looks to be even better in terms of current carrying capacity but is extremely brittle and - given the magnetic forces that would be involved - does not appear to be a realistic possibility.  Apart from the considerable problem of finding a suitable superconducting material, there are also major problems to be faced in the design of the hollow-center dewar.    Employing 3He polarization analyzers near high field sample magnets would thus appear to be a far-from-trivial undertaking.

(iv)  On-line optical pumping

There are two potentially viable approaches to maintaining constant (or near-constant) gas polarization in 3He cells:  in-situ optical pumping and batch gas transfer from an externally pumped cell to the working cell.

Continuous in-situ optical pumping using the spin exchange optical pumping (SEOP) method does not appear to be particularly difficult in the low sample field configuration of Figure 1.

Heating the cell to about 200 C would be necessary to maintain the Rb vapor pressure needed for efficient Rb-3He spin exchange, but this does not look to be an insurmountable obstacle.  For a PASTIS-type arrangement of three orthogonal Helmholtz coil pairs, however, continuous in-situ pumping would be more of a challenge because the circularly polarized pumping light has to enter the cell parallel to the holding field which, in this case, would shift from one direction to another depending on what is regarded as the most experimentally advantageous sample field orientation.

Batch gas transfer from an externally pumped cell connected via capillaries to the working cell also looks to be a viable option in the low sample field configuration of Figure 1.  In this case the approach would be a closed loop arrangement involving either metastable exchange optical pumping (MEOP) or SEOP.  MEOP is the faster of the two methods but the gas has to be optically pumped at low pressure and then compressed and stored under pressure for periodic exchange with gas in the working cell.  At the present time MEOP appears to be the less attractive alternative for batch transfer in part because compressors that can compress the gas without significant loss of polarizationare are still in the development stage and in part because they are almost certain to be costly and require maintenance on a regular schedule.  SEOP, on the other hand, although slower, is not constrained to low pressures and polarized gas exchange could probably be accomplished without any mechanical components other than the valves controlling the gas flow and would be comparatively maintenance-free.  But whatever the choice of pumping method, issues relating to providing the necessary holding field uniformity in the connecting capillaries would need to be addressed.

Given the more demanding nature of the intermediate and high sample field cases, it is likely that 3He polarization analyzers, if employed, will, at least initially, be optically pumped off-line and manually interchanged as the need arises.  On-line optical pumping by either the in-situ or batch method, although not ruled-out, would require a significant developmental effort because of the limited access to cells inside either solenoids or hollow-center dewars.

References:

(1) J.A. Stride, K.A. Andersen, A.P. Murani, H. Mutka, H. Schober and J.R. Stewart - Physica B356, 146 (2005)

(2) J. Dreyer, L.P. Regnault, E. Bourgeat-Lami, E. Lelievre-Berna, S. Pujol, F. Thomas,

M. Thomas and F. Tasset - Nuclear Instruemnts and Methods A449, 638 (2000)