Deterioration of the calculated HYSPEC performance (non-polarized neutron beam) due to several “real world” decisions

Vinita J. Ghosh

May 15, 2006

- HYSPEC moderator-monochromator distance increased from 25 to 37m (straight guides) to accommodate the NSE instrument in the target building.

| Ei(meV) | LMM=25 | LMM=35 | LMM=40 |
| --- | --- | --- | --- |
| 3.6 | 2.2% | 2.2% | 2.1% |
| 15 | 4.5% | 4.3% | 4.2% |
| 30 | 6.4% | 6.0% | 5.9% |
| 60 | 9.0% | 8.2% | 8.2% |
| 90 | 10.4% | 10.1% | 9.8% |

Flux on sample and energy resolution (ref 1): Flux on sample decreases from 1x107  to  6x106 (n/cm2.s) for 15meV, and from 1.5e6   to  8x105 (n/cm2.s) for 60meV.

- The HYSPEC guide was curved to reduce background at the detector bank. The guide curvature was designed to get an offset of 16cm (4 times the guide width) at the focusing crystal (ref 2).

| Nominal Energy Ei (meV) | Fluence(n/s) for offset=0cm | Fluence(n/s) for offset=16cm | Ratio of fluence at the end of the curved guide |
| --- | --- | --- | --- |
| 3.6 | 5.27e9 | 5.26e9 | ~1 |
| 5.0 | 6.24e9 | 6.24e9 | ~1 |
| 15.0 | 4.84e9 | 4.75e9 | 0.98 |
| 30.0 | 9.60e8 | 9.00e8 | 0.94 |
| 60.0 | 2.98e8 | 2.40e8 | 0.81 |
| 90.0 | 1.86e8 | 1.06e8 | 0.57 |

Ref 2, table 4. Intensity as a function of offset for LMM=37m.

- Coating of the inner surface of the curved guide to be m=1. This will result in a significant reduction in flux on sample at the lower energies (ref 2,3).

Ref 3, Fig 5b.

- Chopper Box A containing choppers T0 and T1 moved from 6.5 to 8.5m to allow for better crane coverage and it was tacitly assumed that the curved guide would start at 9.4m. This reduces the length of the curved guide from 24.8 to 23.3m, and the radius of curvature of the guide G2 has to changed from 2542 to 2323m to get an offset of 16cm at the focusing crystal. There is a third (better) option available. The curved guide should start at 6.32m, immediately downstream of the 4m expanding section. The curved guide is now longer (26.38m) and the radius of curvature can be increased to 2884m. However, a gap is required in the curved guide to accommodate the choppers T0 and T1A. Even after allowing for this gap there is an improvement in flux compared to the 9.4m option. The McStas results were written up in ref 3.

Fig 4b. Ratio of the fluence at the end of guide G3 for the three different starting positions of the curved guide: G2start_7.9m, G2start_9.4m and G2start_6.32m with respect to G2start_7.9m.

- Width of the PSI PG focusing crystal array will be 260mm NOT the 300 mm required to completely cover the 40mm wide guide at 90meV. This 260mm crystal array will completely cover the 40mm wide guide only upto 76.8 meV. At 90meV the neutron beam intensity will decrease by a factor of 26/30.

- A straight slotted Fermi chopper will be used as the wavelength selector instead of a counter-rotating double disk chopper. The slot length of the AlGd slit package is expected to be 10mm. The transmission through 10mm of Al can be estimated from the figure below, and it is typically 85%.

Summary/conclusion

Each of the above decisions appear to be very reasonable and the deterioration in performance due to each change in design is relatively small, but the total impact on the performance of HYSPEC (namely on the flux on sample) is substantial. Let R(1) represent the reduction in the flux on sample due to decision 1, and R(NP) the total reduction in flux due to all 6 decisions.

| Energy | R(1) | R(2) | R(3) | R(4) | R(5) | R(6) | R(NP) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 15meV | 0.6 | ~1 | 0.72 | 0.92 | ~1 | 0.83 | 0.33 |
| 90meV | 0.53 | 0.57 | ~1 | 0.88 | 0.87 | 0.87 | 0.20 |

The HYSPEC flux on sample for a nonpolarized neutron beam will be lower than that of the CNCS, ARCS and SEQUOIA at all incident energies (see comparison in ref 4).

HYSPEC flux on sample as planned

Should we take steps to recover the performance of HYSPEC before it is too late?

What can we do to improve the performance of HYSPEC?

Some of the decisions listed above are non-negotiable. This includes going outside the target hall(1), an offset of 16cm at the focusing crystal(2) and the use of a straight slotted Fermi chopper with a AlGd slit package(6). We can change the coating of the inner surface of the curved guide from m=1 to m=2 (3). However, this involves additional expense. We can change the length, Rc and starting position of the curved guide (4) at no additional cost. We can ask PSI to increase the width of the PG monochromator from 260 to 300 mm (5). This should not entail a huge amount of re-engineering and they may respond favorably to our request.

| Energy | R(1) | R(2) | R(3) | R(4) | R(5) | R(6) | R(NP) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 15meV | 0.6 | ~1 | 0.95 | 0.95 | ~1 | 0.83 | 0.45 |
| 90meV | 0.53 | 0.57 | ~1 | 1.14 | ~1 | 0.87 | 0.30 |

These three changes should result in some improvement, though not as much as I had hoped for. At the very least HYSPEC will have a flux on sample as good as the CNCS.

References:

[1] V.J. Ghosh presentation at the HYSPEC IDT meeting (Sept.9, 2004

A comparison of the options for the location of the HYSPEC instrument, V.J. Ghosh, M.E. Hagen, L. Passell, S.M. Shapiro and I. Zaliznyak, (April13, 2004)

[2] McStas simulations for guide configuration optimization, HYSPEC Technical Note 0009, V.J. Ghosh and M. Hagen, (February 2005).

[3] More McStas simulations for the HYSPEC guides. Vinita J. Ghosh, (September 2005).

[4] HYSPEC: A crystal time-of-flight hybrid spectrometer for the SNS, S.M. Shapiro, I. Zaliznyak, L. Passell and V. Ghosh, ICANS-XVI (May 12-15, 2003).