## Slide 1: HYSPEC
Recent progress and polarization analysis capabilities
ORNL: B. Winn, V. Ovidiu Garlea, M. Graves-Brook, Peter Jiang, X. (Tony) Tong
BNL: L. Passell, S.M. Shapiro, I. Zaliznyak
PSI: U. Filges, Michel Kenzelmann
ESS: M. Hagen (formerly ORNL)
WINS 2014, May 2014

## Slide 2: Hybrid Spectrometer: a cross between a direct geometry spectrometer…
Measures variable Ef of scattered neutrons using ToF
L2=3.6 m, L3=4.5m
Short strait blade Fermi chopper at L1=37.2 m to select Ei, and trade off between E resolution and flux via frequency range 60-420 Hz

## Slide 3: …and a Triple Axis Spectrometer
Variable Direction Ar flight path for scattered beam
60o x 15o, r = 4.5 m
Scattered beam
Vertical focusing array: choice between PG and Heusler
Incident beam
Configurable region about sample

## Slide 4: Intensity at sample, V scatter
15 cm high beam from guide vertically focused to sample ~2.5 cm x 2.5 cm FWHM
4.2E5 n/s/MW/cm2: Gold foil measurement at sample position, PG focus array to sample 1.8 m, Ei=15 meV, Fermi 180 Hz
Plot: Vanadium incoherent isotropic scatter integrated over detector array at 40o < 2Q < 100o, PG & Heusler
3.8 meV < Ei < 50 meV
Common Ei’s:
3.8, 7, 15, 20, 27, 35 meV
Rare Ei: 50 meV
Chart: Incident energy (meV)

## Slide 5: Active in User Program, Unpolarized
1 W. Tian et al., Phys. Rev. B 89, 144417 (2014)
2D. Fobes et al., Phys. Rev. Lett, accepted
 | #
 | Experiments
Spring 2013 | 6
Fall 2013 | 17
Spring 2014 | 16
Y0.7Lu0.3MnO3 at 4K: dotted lines show calculated spin wave dispersion with magnetoelectric coupling1
Fe1.09(1)Te: :Bragg peaks visible at 5 K indicate increased Fe displacements from high-symmetry positions in the a-b plane2

## Slide 6: More Science
Spin waves in
S=1/2 square lattice
antiferromagnet K2V3O81
1 M. Lumsden, A. Christianson, in preparation,
Ei=7 meV 180 Hz s2=-45o +/-0.1 meV, L integrated. s1~90o range 0.5o step
2O. Delaire et al, in preparation, TO branch softens on cooling from 300 to 120 K, approaching the Curie temperature for the ferroelectic transition at ~80 K
TO branch softening in KTa0.88Nb0.12O32

## Slide 7: Fits into HFIR/SNS capabilities at Oak Ridge National Lab
(Green: Supports Polarization Analysis )
ARCS
Cold Moderator / 2 meV < Ei < 100 meV
HYSPEC
Thermal Moderator / 5 meV < Ei < 2000 meV
1-3% elastic E resolution / less flux
CNCS
SEQUOIA
3-5% elastic E resolution / high flux
CTAX
HB1A, HB1, HB3
Highest time-averaged flux, limited solid angle
Direct Geometry Spectrometers
Triple- Axis Spectrometers
Stone et al, Rev. Sci. Instrum. 85 p 045113 (2014)

## Slide 8: Polarization Capabilities
Incident beam
Scattered beam
Sample
Polarizer option 1: Heusler focusing array
Magnetic Guide Field assemblies
Flipper option 1: Mezei
Typical Configuration
Heusler focusing element, Mezei flipper, Polarizing supermirror array at sample, 15 meV, best flipping ratio 23
Nonmagnetic materials about sample. Mostly aluminum, titanium and stainless steel
3He cell, Heusler single crystal and polarizing supermirror array for characterization of polarized beam

## Slide 9: Alternatives, polarized beam
Polarizer option 2: in situ Spin Exchanged Optical Pumped 3He cell1
Polarizer option 3: Polarizing supermirror array, either reflection or transmission
PG focusing element for unpolarized beam
Flipper Option 2: High Tc Cryo-Flipper 2
Polarized beam
Polarized beam
Unpolarized beam
Demonstration only. Considering possible flux advantage over Heusler focus array (~3x at 50 meV).
~13.2 flip with supermirror at sample and 15 meV
~3.9 with Heusler at sample and 25 meV
~2.7 with Heusler at sample and 50 meV
Demonstration only. Considering development for use with cryo-magnets.
32.7 system-wide flip ratio with supermirror at sample, 15 meV
2T. Wang et al, Neutron spin manipulation devices using YBCO films, IOP conference series accepted from the Neutron optics, polarisation and detectors meeting in Munich 2013
1X. Tong et al, Rev. Sci. Instrum, 83, 075101
Polarized beam

## Slide 10: Half Polarized Experiments
Sample
Vertical bore, single crystal and powder
Horizontal field permanent magnet prototype. Flexibility during polarization commissioning
Rb2Fe2O(AsO4)2

## Slide 11: Polarization Analysis option 1: 3He
Incident beam
Scattered beam
Sample
Ultimate goal: wide-angle quartz cell. Same depth for all scatter angles
Scattered beam
Polarization analyzer: 3He cell optically pumped via spin exchange (SEOP)
Abberation-corrected Helmholtz-like coils, vertical bore, uniform field centered on analyzer cell, also surrounding sample
Cannot directly SEOP, so:
Drop in cells now
Automated transfer
Analyzed beam

## Slide 12: Drop-in-cell Polarization Analysis
Adiabatic Fast Passage coil to flip 3He
‘banana’ cell made at NCNR
Cylindrical cell, limited solid angle
Lifetimes up to 100 hr
Direct beam transmission flip ratios
Through pinhole at sample
At 3.8 meV: unfocused: 16.3
focused: 21.8
Vanadium scatter measured ratio 1.47

## Slide 13: Auto-Refill, Vary Pressure
Rev. Sci. Instrum. 84, 065108 (2013), C. Y. Jiang, et al.

## Slide 14: Polarization analysis #2: PSI Polarizing Supermirror Array
Scattered beam
Remove radial collimator
- 0.5 meV < Ef < 20 meV
+ Larger sample environments
+ Larger magnetic fields at sample
Analyzed beam
Scattered beam

## Slide 15: Performance and Preparations
960 supermirrors, 60o
U. Filges
BOA beamline at SINQ
Ready to ship
New magnetization unit at HYSPEC
Chart: White beam 2-6 Angstrom
| 25 | 20 | 15 | 10 | 5 | 0 | -5 | -10 | -15 | -20 | -25
Center | 94.8 | 94.1 | 94.8 | 95.1 | 94.5 | 93.4 | 94.5 | 94.4 | 93.4 | 94.3 | 94.8
5 cm above | 92.7 | 94.1 | 93.7 | 93.3 | 93.4 | 91.7 | 94.7 | 95 | 93.5 | 92.9 | 94.1
Chart: White beam 2-6 Angstrom
| 25 | 20 | 15 | 10 | 5 | 0 | -5 | -10 | -15 | -20 | -25
center, optimum polarization | 30.2 | 32.4 | 31.2 | 31.6 | 31 | 29.9 | 30.8 | 34.1 | 29.7 | 32.9 | 30.1
5 cm up, optimum polarization | 33.8 | 30 | 30.1 | 32.5 | 31.3 | 24.5 | 28.6 | 26.3 | 24.7 | 30 | 28.1
optimum flux, center | 40.7 | 36.6 | 41.3 | 40.9 | 40.3 | 41.2 | 44.9 | 45.9 | 42.4 | 43 | 40.9
optimum flux 5 cm up | 40.3 | 35.2 | 38 | 37.5 | 37.5 | 36 | 40.7 | 41.6 | 35.9 | 37 | 39.3

## Slide 16: HYSPEC History & Team
Inception: BNL1
Design and Development2:
  Instrument Development Team
    BNL: S. M. Shapiro, I. A. Zaliznyak, L. Passell, V. J. Ghosh, W. J. Leonhardt
    PSI: U. Filges (polarizing supermirror array analyzer)
  ORNL: M. E. Hagen, D. Anderson, T. Tong
    Many support teams
Install & Unpolarized Commissioning:
  ORNL: MEH, M.Graves-Brook, B. Winn
  Significant IDT & ORNL staff support
User Program & Polarization Commissioning:
  ORNL: MGB, BW, O. Garlea, T. Tong, P. Jiang, D. Brown
[1] I. Zaliznyak, V. Ghosh, S. M. Shapiro, L. Passell, Physica B 356, 150 (2005).
[2] S. M. Shapiro, I. A. Zaliznyak, L. Passell, V. J. Ghosh, W. J. Leonhardt, M. E. Hagen, Physica B, 385-386, 1107 (2006).