## Slide 1: Fly and trajectory scans
Tim Mooney
2/27/2015

## Slide 2: Fly-scan choices
Software fly scan
  Data acquired while positioners move at constant speed
  Detectors triggered by software
    Periodically, or at user-specified time intervals
  Positions acquired by software along with detector data
  Few-ms dead time between data points
  No cabling required
Hardware fly scan
  Data acquired while positioners move
  Detectors triggered by pulses from positioner
    Periodically
    At user-specified positions
  Positions implied or acquired by multichannel scaler
    Arraycalc “cum” function reconstructs motor positions from scaler data
  No dead time between points
  Need cable from selected motor to selected detector(s)

## Slide 3: Hardware fly choices
Constant speed (from now on, “Hardware fly scan”)
  Positioner moves at constant speed
  Positioner can be moved by motor record
  Specify StartPos, EndPos, NPTS, and Speed
  May be able to specify data-acquisition positions
    Requires supported motor or external hardware
Trajectory (from now on, “Trajectory scan”)
  Motor moves along specified trajectory
  Use trajectory controls
  Specify trajectory positions and times
  Specify data-acquisition positions
    Details depend on controller
  Requires supported motor

## Slide 4: Software fly scan
Requirements:
  positioner speed can be set
  position updates periodically or on demand
  scan manager (e.g., sscan record)
User interface:
  Differences from step scan

## Slide 5: Software fly scan
Remarks:
  Positioner speed and detector-dwell time must be set
  Imprecise synchronization between data and recorded positions
  Limited to <~ 10 Hz
PVs: (values in green: user’s choice)
Modified for software fly scan:
$(scan).P1SM = FLY
$(motor).VELO = 1.0
Same as for step scan:
    $(scan).ACQT = SCALAR
    $(scan).P1PV = $(motor).VAL
    $(scan).R1PV = $(motor).RBV
    $(scan).T1PV = $(scaler).CNT
    $(scan).DnnPV = scaler.T
    $(scan).P1SP = 0
    $(scan).P1EP = 5
    $(scan).NPTS = 10
    $(scaler).TP = .5

## Slide 6: Hardware fly scan requirements
Positioner can output pulses during nontrajectory move
  Any stepper motor (with external divide-by-N)
  Aerotech Ensemble with EnsemblePSOFly database
  Probably other servo motors can do this
  softGlue can do this if you give it quadrature-encoder signals
Hardware-triggered detector can cache or stream scan data
  MCS (Struck multichannel scaler)
  XIA DXP
  Cameras supported by areaDetector
Data-storage/management client
  sscan record (saveData)
  Spec, BlueSky
  areaDetector plugin

## Slide 7: Hardware fly scan
User interface:
  Differences from software fly scan:

## Slide 8: Hardware fly scan
Remarks:
  positioner speed must be set
  detector must be prepared and started before motor moves
PVs:
Modified for hardware fly scan:
    $(scan).ACQT = 1D ARRAY
$(scan).BSPV = $(mcs):EraseStart
$(scan).BSWAIT = NoWait
$(scan).A1PV = $(mcs):StopAll
$(mcs):PresetReal = 0
$(mcs):ChannelAdvance = External
$(mcs):Channel1Source = Int. clock
$(scan).D01PV = $(mcs):mca1.VAL
$(mcs):Prescale = 2500
$(mcs):CountOnStart = Yes
$(mcs):NuseAll = 1000
    $(scan).R1PV = not used
    $(scan).T1PV = not used
Same as for software fly scan:
    $(scan).P1PV = $(motor).VAL
    $(scan).P1SM = FLY
    $(scan).P1SP = 0
    $(scan).P1EP = 5
    $(scan).NPTS = 10
    $(motor).VELO = 1.0

## Slide 9: Trajectory scan requirements
Controller can move motor along trajectory
  Newport MM4005 or XPS
  Aerotech Ensemble
  Galil
Controller can generate position-synchronized pulses
Position-table generator
  E.g., spec, arraycalc, python, etc.
Hardware-triggered detector can cache or stream scan data
  Same as hardware fly scan
Data-storage client
  Same as hardware fly scan

## Slide 10: Trajectory scan
User interface:
  Differences from hardware fly scan:

## Slide 11: Trajectory scan
Remarks:
  Trajectory must be loaded
  Detector must be prepared and started before motor moves
PVs:
Modified for trajectory scan:
    $(scan).T1PV = $(traj):Execute
$(mcs):Prescale = 1
$(scan).BSWAIT = Wait
$(scan).BSPV = prepForTraj
$(scan).A1PV = prepData
$(mcs):CountOnStart = No
$(traj):* = many choices
    $(scan).P1PV = not used
    $(motor).VELO = not used
    $(scan).P1SP = not used
    $(scan).P1EP = not used
    $(scan).P1SM = not used
Same as for hardware fly scan:
    $(scan).ACQT = 1D ARRAY
    $(mcs):PresetReal = 0
    $(mcs):ChannelAdvance = External
    $(mcs):Channel1Source = Int. clock
    $(scan).NPTS = 10
    $(scan).D01PV = $(mcs):mca1.VAL
    $(mcs):NuseAll = 1000
    $(scan).R1PV = not used

## Slide 12: Trajectory definition
Number of trajectory elements
Array of positions for each motor
  Ensemble: only one motor
Array of times
  Can be specified as total time
Number of output pulses, start/end element
  MM4005: pulses evenly spaced in distance along trajectory
  XPS: pulses evenly spaced in time
  Ensemble: pulses evenly spaced in distance, or at trajectory points
    Under development: at user-specified positions
  For MAXv: pulses only at trajectory points
Absolute/Relative/Hybrid position mode
  Currently, Ensemble and MAXv don’t support Hybrid mode
MAXv has timing problems in very slow motion

## Slide 13: Detector-trigger options
motor/encoder
MCS
÷N
detector
motor/encoder
MCS
÷N(i)
detector
motor/encoder
softGlue
÷N(i)
detector
MCS
e.g., step-motor hardware fly
e.g., trajectory, ensemblePSOFly
e.g., tableFly
÷1
÷1

## Slide 14: Examples
1ide hexFly (hard fly scan)
  EnsemblePSOfly.db with evenly spaced data-gate signals
2bmb, 32idc tomography fly (hard fly scan)
  EnsemblePSOfly.db with evenly spaced data-trigger signals
2bmb interlace fly (hard fly scan)
  EnsemblePSOfly.db with user-specified data-trigger signals
    tableFly.db uses softGlue to generate triggers from motor pulses
    interlaceFly.db programs tableFly
  Acquire at ~100 Hz for ~30 minutes
15idd USAXS fly (trajectory scan)
  Ensemble trajectories with user-specified data-trigger signals
  Rotation stage executes exponential trajectory
  Translation stages execute commensurate trajectories
Gradient-multilayer-deposition system (trajectory motion)
  Galil trajectory